# Auditable Data Analysis and Management System for ELISA (ADAMSEL-v1.1)

# Conversion of OD into concentrations by a Four-Parameter Logistic Curve Fit

© EJ Remarque 2007

## Contents

Introduction	3

How to use ADAMSEL	3

Disclaimer	4

General conventions in the application	4

Define locations, dilutions and read the OD data	4

Define the plate layout	4

Read the OD data	5

Adding sample names	5

Fit the regression curve	5

Save the results	6

The Settings sheet	10

The Format sheet	11

Merging calculated results from multiple plates	12

## Introduction

ADAMSEL FPL is an application that converts OD readings obtained from ELISA plate readers into concentrations by four-parameter fitting and ADAMSEL Merge is an application to merge results generated by ADAMSEL FPL.  They are designed to provide an auditable system that minimizes data handling, thereby reducing the chances for error.

The equation used is:

OD=MaxOD−MinOD1+elnXmid−lnConcScale+MinOD

where:

MaxOD 	=	Upper asymptote for OD

MinOD	=	Lower asymptote for OD (blank value)

Conc		=	The concentration (or reciprocal dilution)

Xmid		=	The midpoint of the sigmoid curve

Scale		=	The slope of the regression line

The application fits the sigmoid curve in two steps:

- The starting values for MaxOD, Xmid and Scale are estimated by linear regression on logit-transformed data, where the R2 is optimized by changing the MaxOD value. The starting values for Scale and Xmid are then estimated from the linear regression. The MinOD is set at the blank value observed in the data.

- The fit is finalised by iterations of MaxOD, Xmid and Scale; where the Mean Square is minimized by changing MaxOD, Xmid and Scale.

Once the regression parameters are known, the unknown samples can be calculated with the following equation:

Conc=Xmid−ScalelnMaxOD−MinODOD−MinOD−1

## How to use ADAMSEL

Three steps are required to convert OD values into concentrations:

- Definition of the locations and dilutions of standards and blanks and subsequent reading of the OD data. This is all done with the buttons on the Raw Data sheet.

- Fitting the regression curve. The OD data used for the regression can be optimised by exclusion of outlying data points.

- Saving calculated data to a file.

## Disclaimer

ADAMSEL is free software for non-commercial users and comes with ABSOLUTELY NO WARRANTY. Please do not redistribute the application, instead refer potential users to the author of the application (remarque@bprc.nl) as this enables the distribution and maintenance of updated versions for all users.

## General conventions in the application

Areas with a light green background color can be changed by the user. All other regions on the worksheet are deliberately locked and cannot be changed by the user. Do not attempt to paste data into the application as this is blocked to avoid the possibility of introduced error. The current application will only work correctly under Windows™ (MAC OS is not yet supported). Set the decimal separator to a point “.”. Users employing “Continental” settings (Dutch, German, French etc.) need to change the international settings accordingly in the Windows™ control panel. Macro security settings in excel need to be set to medium (Tools -> Macro -> Medium), else the application can not be used.

## Define locations, dilutions and read the OD data

Figure 1. The “Raw Data” worksheet.

### Define the plate layout

Before reading OD data, the plate layout must be defined. The application then uses this layout to transport standards and blanks to the “Regression” worksheet. The location of standards, blanks and samples is indicated by selecting the appropriate range in the plate layout on the “Raw Data” sheet (green background area) and subsequently pressing the appropriate define buttons for Standard, Blank or Sample. When defining standards or samples, dialog boxes are used to specify dilutions and dilution factors. For the standard, an additional dialog box is used to specify the highest concentration. Dilutions belonging to standards and samples will automatically be displayed on the “Raw Data” sheet.

Standards are indicated by a name followed by an index letter and a number. A capital index letter (A or B) defines the standard duplicates, and the index number (two digits) defines the sequential dilution of the standards. The standard can maximally use 2 x 12 wells. For example StdA01, indicates the first dilution of the first standard replicate.

Blanks are indicated by a name and an identifier number; up to 96 blanks can be used. For example Blank01, indicates the first blank

Samples can be single or duplicate; they are indicated by S_ followed by the sample number and a capital character to indicate the order. A second lower case character (a or b) is added to indicate sample duplicates. Samples can have up to 12 wells and can be either single or duplicate. For example S01Aa and S01Ba, indicate the first and second dilution of the first replicate of sample 1, respectively.

Formats can be saved and loaded with the buttons provided. The name of the format file is displayed on the “Raw Data” sheet upon opening or saving.

### Read the OD data

OD data from a text (ASCII) file can be loaded by pressing the “Open” button. A file open dialog will appear and the file can be selected. OD input files can be either tab delimited text files or comma-separated files. The name of the file opened is automatically displayed on the “Raw Data” sheet. The application will automatically convert data files containing commas as a decimal separator. Overflow and non-numeric values in the OD data file will be displayed as 9.999 after the OD data have been read. The Standard and blank data will automatically be transferred to the “Regression” worksheet.

## Adding sample names

After the OD data have been read, names can be read from a text file or names can be manually entered on the “Format” sheet. Do not enter names before opening OD data, as names will be automatically reset to default values when an OD data file is opened.

Fit the regression curve

The first fit provided on the “Regression” worksheet contains the start values for the regression, based on a linear fit on logit-transformed OD data. Data points can be excluded or included by ticking the check boxes next to the values. Overflow as well as values below the low cut-off are automatically excluded and can not be selected (the lower cut-off can be defined on the settings sheet). The start values will change accordingly. In a similar way, blanks returning values that are obviously out of range can also be taken out of the calculation, by deleting the outlier value. The start values will then automatically be recalculated. The fit must be finalized by pressing the ”Fit Curve” button. This starts the minimisation of the Mean Square. A report on the optimisation procedure is presented on the “Regression” worksheet. The fit can be repeated after standard points have been in- or excluded. The standard and blank values displayed on the “Regression” sheet can be reset to the start values by pressing the “Initialise Data” button. Comments can be entered, viewed and edited by pressing the “Add comment” button. Comments will be displayed on the “Results” sheet.

After the curve has been fit, the results will be automatically displayed on the “Results” sheet. When the standard changes, either by ticking check boxes or by deleting blank values, the results sheet will be cleared and the final fit must be repeated.

Figure 2. The “Regression” worksheet.

## Save the results

Go to the “Results” sheet and save its contents by pressing CTRL + s (saves as a text file with extension ELI) or CTRL + e (saves as a write protected Excel sheet with extension XLE). A file save dialog box will appear prompting with the filename of the OD data file. The results can be printed by pressing CTRL + p. This will print all pages of the results. Figure 3 shows the first page of the results sheets. The name of the OD data file is displayed on all pages of the results. A regression report is supplied showing the following values:

Figure 3. First page of the “Results” worksheet.

A plot showing the observed standard values (circles and triangles) as well as the predicted OD values (line) is displayed next to the regression parameters.

The standard values are tabulated under the graph. Standard @ and dilution respectively indicate the concentration and dilution of the standard in that particular well. The OD values, calculated concentrations and their corresponding CV values are displayed for each (pair) of observation(s). Incl1 and Incl2 indicate whether the well was included (T for True) or excluded (F for False). The values for the blanks are displayed immediately below the standard values. If a blank value has been deleted on the regression sheet it will be displayed with a strikethrough as well as with an “F” under the heading blank included. Included blank values are indicated by a “T”.

| Field | Description |
| --- | --- |
| Operator | Initials of the regression operator |
| Date | Date when regression was performed |
| Rsquare | Correlation coefficient of observed and predicted OD values. The Rsquare is calculated on log-transformed OD values |
| Mean Sq x 1000 | The value of the mean square multiplied by 1000; defined as: (Observed - Predicted)2/Predicted divided by the number of observations included |
| MaxOD | The Plateau OD value estimated by the regression procedure |
| Scale | A parameter indicating the steepness of the curve, where larger values indicate flatter curves. |
| Xmid | The midpoint of the sigmoid curve |
| Blank | The lower asymptote of the regression curve. This is in fact the average blank value. |
| Blank CV | The coefficient of variation of the blank values |
| High cut off OD | Concentrations for OD values above this value will be displayed as “High” |
| Low cut off OD | Concentrations for OD values below this value will be displayed as “Low” |
| Calculated [Std] | The concentration calculated for the standard using the curve fit parameters. This is the standard regressed on itself |
| N Std. Incl. | The number of data points included for the standard |
| OD = 1 + blank @ | Indicates at which concentration an OD of 1 over blank will be achieved |

Figure 4. Second page of the “Results” worksheet

The results for the individual wells are displayed on page 2 of the results worksheet. The sample names (if entered), dilutions, OD values and concentrations as well as the CV values are presented in a table.

Figure 5. Third page of the “Results” worksheet

The third page of the “Results” sheet shows the aggregated data for each sample. Where a sample has multiple dilutions, the data are reduced to one value for each sample. Duplicates with one or both concentrations at Low or High are excluded from the calculations, as well as samples with a CV higher than the specified maximum CV for duplicates (Specified on the “Settings” sheet). The number of data points included as well as the minimum and maximum OD values for the included data points are tabulated. Concentrations are displayed as the average of all included data points as well as a weighted average. Weights are assigned based on OD values (with blank subtracted), where OD values over 2.000 are assigned a weight of 1; OD values between 2.000 and 0.500 are assigned a weight of 100; OD values smaller then 0.500 and greater then 0.200  are assigned a weight of 10 and OD value lower or equal then 0.200 are assigned weight 1.  This weighting is designed to make concentration estimates for unknows more dependent on OD values within the linear range of the standard curve. The CV’s for the averaged concentration as well as for the weighted concentration are also displayed. The aggregated data are saved in the results file as well as in a separate file with the same name, but with extension “uni” (for units). The “eli” and “xle” files can then later be merged to obtain a file with all titration results that enables auditing (see later under ADAMSEL Merge).

## The Settings sheet

Figure 6. The “Settings” worksheet

The settings worksheet enables the user to change seven application settings. Only values displayed on a green background can be changed by the user.

| Field | Description |
| --- | --- |
| # empty lines in OD file | Some plate readers add additional lines in their output before the actual OD data are given. The number of lines specified here will be skipped when reading OD data from a file. |
| Overflow value | OD readings greater than the value specified will be displayed as overflow (9.999) on the “Raw Data” sheet these values will be ignored by the application and a “High” value will be returned for the corresponding concentrations |
| File Extension for OD File | This can be used to filter the files shown in the OD file open dialog box. In order to avoid seeing too many files, specify the file extension for the OD data files in this field |
| File Extension for Names File | This can be used to filter the files shown in the Names file open dialog box. In order to avoid seeing too many files, specify the file extension for the Names files in this field |
| Tolerance High OD | The value specified here will be used for extrapolations. The value will be added to the maximum OD value observed for the standard. Concentrations for OD values above the highest standard OD observed plus the specified value will be returned as “High” |
| Tolerance Low OD | The value specified here will be used for the low cut off. Concentrations for OD values below blank value * tolerance OD low will be returned as “Low |
| CV cut off | Duplicates with a CV value higher than the one specified here will be excluded from the calculation of the aggregated means |

## The Format sheet

Figure 7. The “Format” sheet

The format sheet should be used to add sample names after OD data have been read. Note that sample names are automatically reset to default values (Sample01 etc.) when OD data are read. The “Format” sheet also offers the opportunity to change sample dilutions and dilution factors. These are initially entered when defining a format, but can be changed later. Any change made on the “Format” sheet will be automatically implemented. If sample dilutions change, the changes are displayed on the “Raw data” and “Results” sheets. Settings for Standards and blanks can not be changed. The standard and blank, can, however, be re-entered after they have been cleared, using the buttons on the “Raw Data” sheet.

## Merging calculated results from multiple plates

Once a series of ELISA plates has been calculated, the results can be merged into one Excel workbook, allowing for easy inspection of the results from whole series. Usually only the calculated data for the samples are saved.  The “ADAMSEL merge” application, however, allows for the aggregation of all results obtained with the ADAMSEL FPL application. The results obtained for Regression, Standards, Blanks, Separate Samples and Aggregated Samples will be displayed on separate worksheets.

Figure 8. Main screen of the ADAMSEL Merge application

After pressing the “Add Files” button, a dialog box will appear and the files to be merged can be selected; only files of the type indicated at File Type will be shown in the file open dialog box. Two values can be used: eli, for text results files and xle for Excel results files. After pressing the “Merge Files” button all results will be merged to 5 separate worksheets within the workbook. The data can then be saved as an Excel workbook by pressing the “Save Results” button. The Excel file thus generated will contain all the results without any Excel Macros. All information contained in the results files will be displayed on the worksheets generated. Strikethrough values on the Standard and Blank worksheets indicate that a certain value was not included in the calculations.

Figure 9. Regression worksheet of the ADAMSEL Merge application

Figure 10. Standards worksheet of the ADAMSEL Merge application

Figure 11. Blanks worksheet of the ADAMSEL Merge application

Figure 12. Samples all worksheet of the ADAMSEL Merge application

Figure 13. Samples worksheet of the ADAMSEL Merge application